Use of autoclave in HS264

To sterilize solids (like flasks, bottles and tubes):

Label items with autoclave tapePlace items in heat-resistant plastic tub(s).

Shut door, making sure Door Closed light comes on and door lock is securely tightened.

Note:  This light will come on before the door is actually secured. Hand tighten door.

Select 20 minutes sterilize time and 0 minutes dry time.  	This cycle generally takes about 45 minutes.

Push SOLID button and make sure pressure and temperature begin to rise.

When cycle is COMPLETE, open door slowly and remove items with heat-resistant gloves.

To sterilize liquids (like media and buffers):

Label items with autoclave tape, in additional to contents, your initials and the date (on regular tape).

Make sure lids are loosely tightened to allow air escape.

Place items in heat-resistant plastic tub(s).

Shut door, making sure Door Closed light comes on and door lock is securely tightened.

Note:  This light will come on before the door is actually secured. Hand tighten door.

Select 15 minutes sterilize time for volumes less than 1 liter to be sterilized.  Add 5 minutes for each additional liter of liquid volume.  Solids may be included in autoclave run.

A 20-minute liquid cycle generally takes about 1 hour.

Push LIQUID button and make sure pressure and temperature begin to rise.

Note:  Exhaust will take longer for the liquid than for the solid cycle.

When cycle is COMPLETE, open door SLOWLY and CAREFULLY remove items with heat-resistant gloves.

Note:  Be very careful moving autoclaved liquids since they can easily boil over.