Protocol for using the Nikon Eclipse 50i Epifluorescence microscope to view GFP

– written by Kim Bowen instructed and modified by P. Erickson 2.17.10

To use the regular scope to find object of interest:

- Turn on Power Button that is on the left hand side of the microscope.

- Adjust intensity of light by the dial on the right hand side of the microscope.

- Match phases – Match phase on bottom dial with the phase that is on the objectives.

- Turn the upper rotational filter dial to the 4, 5, or 6 position.

To use the filters for viewing GFP:

- Focus on the object of interest under white light first.

- Turn on both light sources on the two machines (Xcite and ExFo) on the left hand side of the microscope.

- The shutter (below the filter dial) needs to be in the O (open) position.

- Turn the upper rotational filter knob to 2 position (images will now appear green).

- Turn the lower phase dial to filter to the C (closed) position to block the white light coming up from the bottom of the microscope.

- There should be a blue light shining down onto your slide when you are not looking through the orange filter guard.

- The three filters in the back left side of the microscope (ND4, 8 and 16) are the different options for adjusting intensity and altering signal to background.

When done with looking at the object of interest:

- Turn off all of the power sources

- Put the cover back on the microscope

Using the camera software:

- Turn on the computer

- Click the AdvanceEco Icon

- Press Capture to take Picture of specimen

UV filter (UV-2E/C) in position 1 gives blue images for viewing DAPI, FastBlue, AMC and lipofuscin (we hope!).  Excitation 360 nm and emission 435-485 nm.

Blue filter (B2 E/C) in dial position 2 gives green images for viewing GFP, fluorescein isothiocyanate (FITC), acridine orange, Alexa fluor, rsGFP, FAM, SYTOX Green.  Excitation 465-495 nm and emission 515-555 nm.

Green filter (G2E/C) in dial position 3 gives red images for viewing acridine red, calcium orange, ethidium bromide, Texas Red, TRITC.  Excitation 528-533 nm and emission 590-650 nm.